OTF-1200X-III-3 Split Three-Zone Tube Furnace

with 80 mm dia X 1400 mm L Quartz Tube & Vacuum Flange

# 1. Introduction

OTF-1200X-III-3" is long split-able 3" (I.D)  three-zone tube furnace, which can achieve faster heating up to 1200 oC and create a different thermal gradient by adjust three zone temperature. The furnace includes one 80 mm O.D x 1400mm L fused quartz tube, one pair of vacuum sealing flanges with pressure gauge & ball valve for immediate use.  Three precision temperature controllers are installed, which can provide 30-segment heating and cooling steps with +/- 1 oC accuracy with 20" length constant temperature zone.  It is an excellent furnace for annealing, diffusing and sintering sample in various atmospheres.

# 2. Technical Specifications

| Item | Parameter |
| --- | --- |
| Power | 7KW |
| Voltage | AC 220-240V Single Phase, 50/60 Hz (40A breaker required) |
| Max. Temperature | 1200 oC |
| Continuous Temperature | 1100 oC |
| Max. Heating Rate | ≤ 20 oC /min |
| Suggested Normal Heating Rate | ≤10 oC /min |
| Tube Size and Materials | OD: 80m; ID: 74m; Length: 1400 mm, fused Quartz tube |
| Heating Zone Length | Zone 1: 220mm + Zone 2: 440mm + Zone 3: 220 mm (880mm, 35" in total) |
| Constant Temperature Zone | 20" ( 500 mm) length ( three zone at equal temperature within +/- 1oC) |
| Temperature Controllers | PID automatic control via SCR ( Silicon Controlled Rectifier) power control, 30 steps programmable |
| Temperature Accuracy | ±1 oC |
| Heating Elements | Fe-Cr-Al Alloy doped by Mo |
| Dimension | 1100 mm x 450 mm x 670 mm |
| Warranty | One year but excluding quartz tube |
| Net Weight | 180 lbs |

# 3. Specification

Figure 1

1. Furnace Body

2. Flange

3. Gas In/Out

4. Gas Meter

5. Temperature Controller 1, 2, 3

6. Tube

7. Function Button (ON/OFF/LOCK & Indicating Light for Three Controller)

8. Gas Valve

9. Handle

10. Metal Lock

MTI Corporation                                                            Website: www.mtixtl.com

860 South 19th Street, Richmond, CA 94804, USA

Tel: (510)525-3070     Fax: (510) 525-4705                     E-mail: info@mtixtl.com

# Text boxes

7

3

3

4

2

10

6

8

1

9

5