HPF-800 Split High Pressure Tube Furnace with 80mm SS316 Tube up to 800C @150 PSI

HPF-800 is a split tube furnace with SS316 steel tube (80mm O.D x 70 mm I.D x 1000 mm L) and flanges / valves. It is designed for using at 88 PSI ( 6 atm) high pressure up to 850oC under oxidizing and inert gas atmosphere to process special compound.

Specifications:

| Tube Materials: | Stainless Steel (SS316) |
| --- | --- |
| Tube Size: | 2-1/2" O.D x 1-1/2" I.D. x 40" Length |
| Heating elements | Fe-Cr-Al Alloy doped by Mo |
| Max. Working Temperature: | 800 oC |
| Temperature Controller: | PID automatic control via current limiting phase angle fired the resistor with 30 segments programmable |
| Max. Heating Rate: | <30 oC / min |
| Heating Zone | 440 ( single zone ) |
| Constant Temperature Zone | 200 mm (+/-1 oC) |
| Pressure control | One digital pressure sensor installed on bottom of tube flange and connect to LED display You can use the Pressure controller to set the highest pressure alarm and furnace heating power will be cut-off once pressure reaches the alarm point. |
| Working Pressure VS. temperature | Warning: Must limit Working temperature under a certain pressure limited by vessel material strength 1800 PSI max at 600 oC (124 bar ) 750 PSI max at 650oC (51 bar) 450 PSI max. at 700oC (31 bar) 150 PSI max. at 800 oC (10 bar ) Never use pressure vessel above 800oC |
| Input: | Single Phase, 220V AC, 50/60 Hz |
| Output: | Max. 2.5 KW |
| Dimension: | 550mm x 380mm x 520mm |
| Net Weight: | 150 kg ( 330 lbs ) |
| Dimension of shipment: | 47" x 45"x 45"(1200L mm x 1150W mm x 1150H mm) |
| Weight of shipment: | 200 Lbs |
| Warranty | One year limited warranty for electronic and mechanical part, excluding heating element and ceramic tube damaged by user |

MTI Corporation

860 S 19 St., Richmond, CA  94804, USA

Tel: 510-525-3070     Fax: 510525-4705

E-mail:  info@mtixtl.com     Website:  www.mtixtl.com