Typical Temperature Calibration for GSL1500X-40 Tube Furnace

In order to achieve accurate firing temperature, you shall calibrate temperature between inside tube and outside tube before using by yourselves via NIST certified precision temperature calibrator  (click to order one) and thermocouple.

Please be advised that GSL1500X40 alumina tube furnace is designed for use above 1300oC.  Above 1300oC, temperature inside tube is lower than outside

- The furnace is shipped with temperature off-set. , which is indicated by label on bottom of controller. However, you shall set temperature off-set after calibrating furnace at your desired temperature ( e.g. set SC value in controller, please read manual to learn how )

- The tube furnace‘s constant temperature zone length (+/-1oC) at > 1200oC is about 76 mm in center area.  Please see typical temperature profile below.   You shall place your sample within the zone.

If you have any question for operation the furnace, please feel free to contact us by email at info@mtixtl.com or review our website Tech Support Center to learn

- How to set up ceramic tube and vacuum flange for MTI Tube furnace

- How to calibrate Tube Furnace

- How to use and maintain Alumina tube in MTI high temperature tube furnace

MTI is not responsible for any damage caused by improperly using the furnace

# MTI Corporation

860 South 19th street, Richmond, CA 94804, USA

tel: 510-525-3070, fax: 510-525-4705 email: info@mtixtl.com

www.mtixtl.com

Tech Support Center