How to Check and Replace Heating Element for OTF1200 split-able Tube Furnace

1. Open the furnace as figure 1:

- Unscrew the ten hex-screws on cover (indicated in red No.1-No.10) and then remove the two covers      (indicated in yellow) as figure 2:

3. Following figure 3, respectively, loosen the nut (shown as No.1) and then remove the

power wire (shown as No.2) but keep the resistance wires (shown as No.3) still.

Note: When remove the power wires, please make sure the power is turned off.

4.      Turn the knob pointer of the multimeter (order multimeter here) to Ω scale, and then as figure 4, check resistance between two resistance wire from heating element .  If resistance on millimeter show zero, it means heating element work, otherwise please check connection on heating element, or replace whole heating element with refractory insert    Note: Before measurement, please make sure the nut and the resistance wires well  touch the bolt.

Figure 4

5,  In same way, bottom heating element can be checked and replace

6. If you really need to replace heating elements, please contact MTI at info@mtixtl.com to get further information

# Text boxes

Figure 1

Figure 1

Figure 2

Figure 3